## Slide 1: SPICA the new 6T visible combiner for the CHARA Array
D. Mourard (PI), A. Meilland, P. Berio, J. Dejonghe, D. Lecron,
F. Morand, D. Salabert, F. Allouche, N. Anugu, S. Bosio,
L. Bourges, O. Creevey, S. Deheuvels, A. Domiciano de Souza,
N. Ebrahimkutty, D. Gies, R. Ibanez Bustos, J. Jonak, R. Ligi,
G. Mella, N. Nardetto, K. Perraut, S. Rousseau, M. Vrard,
G. Schaefer, A. Spang, and M. Wittkowski

## Slide 2: SPICA in a nutshell
SPICA-VIS 6T-combiner (0.6-0.9μm)
  3 spectral-dispersions
  spatial filtering with monomode fibers
SPICA-FT: H-band 6T- fringe-tracker
  using a ABCD-combiner integrated-optics
  installed on MIRCX (using the detector and replacing the All-in-One combiner)
MODES | Nb of SpCh | SpCh | Spectral Band | MagLim V²=0.6 | MagLim V²=0.6 + FT | MagLim Vdiff | MagLim Vdiff+FT
LR: R=140 | 60 | ~5nm | 300nm | 8.5 | 11.5 |  | 
MR: R=4400 | 500 | 0.17nm | 85nm |  |  | 5.5 | 9.5
HR: R=13000 | 500 | 0.06nm | 29nm |  |  | 4.5 | 8.5
MagLim: V²: SNR=10, 10mn of integration, for one spectral channel
Vdiff: [SNR=10, sf<5°], V²=0.6 in the reference channel, 30mn of integration, for one spectral channel
SNR calculator based on FRIEND calibration (Martinod+2018), CHARA-AO hypothesis (SR=20%), Vinstr=0.9

## Slide 3: SPICA-VIS installed at Mount Wilson in Feb 2022
Spectrograph Table
Detail of the dispersion unit:
Low Resolution R = 140
Medium Resolution R = 4400
High Resolution R = 13000

## Slide 4: SPICA-VIS installed at Mount Wilson in Feb 2022
General view of the injection table
Pupil Control unit
Injection modules including DDLs
Fast tip/tilt
Image and Pupil plane control
Notes: J’ai rajouté la mention du fast tip/tilt pour la stabilisation de l’injection mais ça fait pas très beau. A toi de voir
J’ai aussi mis la mention en haut à droite… à voir aussi

## Slide 5: SPICA-VIS commissioning in 2023
All 15 fringes simultaneously on SPICA with MIRCx and MYSTIC (June 23)
Implementation of chromatic OPD validated in September 2023
Organisation of night based on ‘survey mode’  one data point every 30mn
End of night automatic pipeline up to L0 and L1 data (uncalibrated  OiDB)
Calibration pipeline validated (work for SPICA, MIRCx and MYSTIC data)

## Slide 6: SPICA-VIS commissioning in 2023
All 15 fringes simultaneously on SPICA with MIRCx and MYSTIC (June 23)
2D power-Spectrum on the 3 instruments showing 15 fringe-peaks on a bright calibrator (HD3360)

## Slide 7: SPICA-VIS commissioning in 2023
All 15 fringes simultaneously on SPICA with MIRCx and MYSTIC (June 23)
Implementation of chromatic OPD validated in September 2023
Correction of the longitudinal dispersion
Common LDC updated every 6s
 small OPD offset in the OPLE control
Visible LDC updated every 6s
SPICA internal DDL adjusted simultaneously
waterfall of SPICA fringes over 10min

## Slide 8: SPICA-VIS commissioning in 2023
All 15 fringes simultaneously on SPICA with MIRCx and MYSTIC (June 23)
Implementation of chromatic OPD validated in September 2023
Organisation of night based on ‘survey mode’  one data point every 30mn
SPICA Night-Scheduler Software
Catalogue from Interferometric Survey of Stellar Parameters (ISSP)
Filtering on criteria (magnitude, declination…)
Adding Calbrators from ISSP cat. And JSDC
Communication with ASPRO2 for the night management
Then ASPRO-A2P2 to SPICA software for OBs
SPICA-DB and Night Scheduling Software
collaboration

## Slide 9: SPICA-VIS commissioning in 2023
All 15 fringes simultaneously on SPICA with MIRCx and MYSTIC (June 23)
Implementation of chromatic OPD validated in September 2023
Organisation of night based on ‘survey mode’  one data point every 30mn
End of night automatic DRS pipeline  L0 (header) and L1 (uncalibrated data)  OiDB

## Slide 10: SPICA-VIS commissioning in 2023
All 15 fringes simultaneously on SPICA with MIRCx and MYSTIC (June 23)
Implementation of chromatic OPD validated in September 2023
Organisation of night based on ‘survey mode’  one data point every 30mn
End of night automatic DRS pipeline  L0 (header) and L1 (uncalibrated data)  OiDB
Calibration pipeline validated (work for SPICA, MIRCx and MYSTIC data)

## Slide 11: SPICA-VIS commissioning in 2023
All 15 fringes simultaneously on SPICA with MIRCx and MYSTIC (June 23)
Implementation of chromatic OPD validated in September 2023
Organisation of night based on ‘survey mode’  one data point every 30mn
End of night automatic DRS pipeline  L0 (header) and L1 (uncalibrated data)  OiDB
Calibration pipeline validated (work for SPICA, MIRCx and MYSTIC data)
However, there are still some issues…

## Slide 12: Injection is still quite low
Strehl Ratio ~ 8%
expected 25%
Loss of flux ~ x10
 2.5mag
No Tel. AO yet!

## Slide 13: Alignment & image quality to be improved
Strong static aberrations
Infrared and Visible co-alignment procedure not yet robust
→ drifts of images through observation
Notes: Tu peux peut-être prendre le fichier gif que j’ai envoyé hier, c’est le plus récent et ça montre au moins les 6 AO en fonctionnement mais les aberrations statiques

## Slide 14: Transfer function on sky is quite-low
a loss in SNR
by a factor 25 compared to the calibration lamp

## Slide 15: Transfer function on sky is quite-low
a loss in SNR
by a factor 25 compared to the calibration lamp
a CALIBRATION LAMP

## Slide 16: Transfer function on sky is quite-low
Important oscillations in the optical delay lines
Investigations under progress
200nm
~l/3.5
a loss in SNR
by a factor 25 compared to the calibration lamp
Notes: Peut-être inutile de rentrer trop dans ce détail à SPIE.

## Slide 17: Statistic on V² measurements on sky
Model fitting
errV2 = 0.05+10 ^(0.7*magCorr-2.75-0.37*mag)
Existence of a plateau in V² accuracy of 0.05
not yet understood!
Might be a detector issue

## Slide 18: First Science observations
From Interferometric Stellar Parameter Survey (ISSP)
: CHARA/SPICA interferometric observations
of a large sample of stars all over the HR diagram
From Stars to Planets
Objective 1 : exoplanet host stars (
Objective 2 : asteroseismic targets (
Objective 3 : calibration of Surface Brightness Color Relations (SBCR)
Stellar evolution
Age
Transit of exoplanets
From Bright to Faint stars
Transit of stars
Objective 4 : limb-darkening over HR diagram
Objective 5 : Stellar activity over the HR diagram: binarity, rotation, wind & environment
Control of systematics
Strong links with:
- ESA/PLATO mission
- Araucaria program
radius, mass and age of planets
distances in Universe (Ho)
The ISSP science team
R. Ligi, M. Vrard, R. Ibañez-Bustos, N. Ebrahimkutty, J. Jonak
CoI of the Survey
O. Creevey (astero-dwarfs),
S. Deheuvels (astero-giants),
A. Domiciano de Souza (rotating stars),
N. Nardetto (SBCR),
M. Wittkowski (winds), and K. Perraut

## Slide 19: First Science observations
Example of stellar diameter measurement
Noise plateau around V²≈ 0.03
S1S2
E1E2
W1W2
HD 222107
mR=3

## Slide 20: First Science observations
The bright supergiant Rigel
Noise plateau around V²≈ 0.01
mR=0!
S1S2
W1W2
S1E2
W2-S2

## Slide 21: First Science observations
Multi-wavelength study of the Binary Star β Tri
Blue = SPICA (R)
Light Pink = MIRCX (H)
Red = MYSTIC (K)
Blue = MIRCX (H)
Red = MYSTIC (K)
No SPICA CP
Simultaneous Data obtained from 0.6 to 2.4µm
Analysis performed by J. Jonak as part of his PhD thesis
Modelling with 2 UDs with chromatic flux ratio
MCMC run to find the best model and estimate parameters and uncertainties
Juraj Jorak
PhD at Lagrange Lab.

## Slide 22: First Science observations
Multi-wavelength study of the Binary Star β Tri seen from 0.6 to 2.4μm
Flux ratios in R, H, K
Diameter 2
Diameter 1
X position
Y position

## Slide 23: What about SPICA-FT?

## Slide 24: SPICA-FT status
SPICA-FT chip test in Nice in 2022
SPICA-FT chip installed in MIRCX in 2024
Chip Installation at CHARA

## Slide 25: SPICA-FT status
SPICA-FT IO chip ~ 30% gain compared to MIRCx in term of transmission
Automatic switch between the MIRCX-AIO and the SPICA-FT-ABCD
10mn for alignment and cophasing
Hardware tests at CHARA
Notes: Nous avons changé les paraboles hors axe de MIRCX afin de tirer profit d'une meilleure adaptation à l'injection en présence d'une bonne correction par l'optique adaptative. Pour faire ce travail nous avons démonté entièrement l'injection MIRCX et donc tout réaligné en ajoutant des étages importants pour le contrôle de la pupille. Cela a permis un gain quantitatif sur le flux injecté.

## Slide 26: SPICA-FT status
  Real Time processing for Group Delay and Phase Delay tracking validated (including relocking fringes)
  Good performance but not enough time to use it in science observations
Software tests at CHARA
Tracking using
Group Delay & Phase Delay
Freq = 300Hz
rms ≈ 100nm

λ / 7 for SPICA-VIS
Notes: Nous avons changé les paraboles hors axe de MIRCX afin de tirer profit d'une meilleure adaptation à l'injection en présence d'une bonne correction par l'optique adaptative. Pour faire ce travail nous avons démonté entièrement l'injection MIRCX et donc tout réaligné en ajoutant des étages importants pour le contrôle de la pupille. Cela a permis un gain quantitatif sur le flux injecté.

## Slide 27: Conclusions
SPICA-VIS:
  Functional commissioning almost completed in 2023
  Routinely work in Survey mode: Scheduler + ASPRO2 & OIDB integration, data reduction ready
  Limited performances on-sky :
    AO, alignment, DLs vibrations  current work by the SPICA & CHARA Teams
    0.02 V² “plateau”  still being investigated  ANDOR detector issue?
SPICA-FT:
  Successful installation on MIRCX (easy switch for the two modes)
  Software ready for GD & PD including fringe jumps (relocking fringes)
  Next steps:
    Test in routine “science” observations for robustness
    Characterisation on SPICA-VIS sensitivity with longer DIT
SPICA the new 6T visible combiner for the CHARA Array
Notes: J’ai enlevé les deux derniers items.